## Sheet1
Kind of uncertainty | Instrument Type | Condition | Tolerance + dB | Tolerance - dB | worst case uncertainty + dB | worst case uncertainty - dB | worst case uncertainty ± dB | standard uncertainty + dB | standard uncertainty - dB | standard uncertainty ± dB
Level error vector signal generator | R&S® BTC |  | 0.5 | 0.5
Additional level error I/Q modulator | R&S® BTC |  | 0.3 | 0.3
Attenuation tolerance of splitter | R&S® RVZ |  | 0.1 | 0.1
Attenuation tolerance of matching pad | R&S® RAM |  | 0.15 | 0.05
Absolute level measurement uncertainty | R&S® NRP6A | between -70 dBm and -20 dBm (+15 °C to +35 °C) | 0.081 | 0.081 | [6 empty] | Zero offset absolute (pW) | 235 | Noise absolute (pW) | 49
Absolute level measurement uncertainty | R&S® NRP6A | between -20 dBm and 0 dBm (+15 °C to +35 °C) | 0.083 | 0.083
| R&S® NRP6A | @ -4 dBm after external zeroing | 2.5636e-06 | 2.5636e-06 | [6 empty] | Signal level (dBm) | -4
| R&S® NRP6A | @ -4 dBm | 5.3454e-07 | 0
| R&S® NRP6A | @ -12 dBm after external zeroing | 1.6175e-05 | 1.6175e-05 | [6 empty] | Signal level (dBm) | -12
| R&S® NRP6A | @ -12 dBm | 3.3727e-06 | 0
Zero offset | R&S® NRP6A | @ -23.3 dBm after external zeroing | 0.00021819 | 0.00021819 | [6 empty] | Signal level (dBm) | -23.3
Measurement noise | R&S® NRP6A | @ -23.3 dBm | 4.5497e-05 | 0
Zero offset | R&S® NRP6A | @ -30 dBm after external zeroing | 0.0010205 | 0.0010205 | [6 empty] | Signal level (dBm) | -30
Measurement noise | R&S® NRP6A | @ -30 dBm | 0.0002128 | 0
Zero offset | R&S® NRP6A | @ -35 dBm after external zeroing | 0.0032262 | 0.0032262 | [6 empty] | Signal level (dBm) | -35
Measurement noise | R&S® NRP6A | @ -35 dBm | 0.00067289 | 0
Zero offset | R&S® NRP6A | @ -42.7 dBm after external zeroing | 0.018963 | 0.018963 | [6 empty] | Signal level (dBm) | -42.7
Measurement noise | R&S® NRP6A | @ -42.7 dBm | 0.0039608 | 0
Zero offset | R&S® NRP6A | @ -50 dBm after external zeroing | 0.10088 | 0.10088 | [6 empty] | Signal level (dBm) | -50
Measurement noise | R&S® NRP6A | @ -50 dBm | 0.021228 | 0
Zero offset | R&S® NRP6A | @ -51 dBm after external zeroing | 0.12662 | 0.12662 | [6 empty] | Signal level (dBm) | -51
Measurement noise | R&S® NRP6A | @ -51 dBm | 0.026708 | 0
Zero offset | R&S® NRP6A | @ -55 dBm after external zeroing | 0.31131 | 0.31131 | [6 empty] | Signal level (dBm) | -55
Measurement noise | R&S® NRP6A | @ -55 dBm | 0.066779 | 0
Zero offset | R&S® NRP6A | @ -58 dBm after external zeroing | 0.60046 | 0.60046 | [6 empty] | Signal level (dBm) | -58
Measurement noise | R&S® NRP6A | @ -58 dBm | 0.13224 | 0
Zero offset | R&S® NRP6A | @ -63 dBm after external zeroing | 1.6699 | 1.6699 | [6 empty] | Signal level (dBm) | -63
Measurement noise | R&S® NRP6A | @ -63 dBm | 0.40511 | 0
Expanded level measurement uncertainty | R&S® NRP6A | @ -4 dBm after external zeroing |  |  | 0.083003 | 0.083003 |  | 0.083 | 0.083
Expanded level measurement uncertainty | R&S® NRP6A | @ -12 dBm after external zeroing |  |  | 0.08302 | 0.083016 |  | 0.083 | 0.083
Expanded level measurement uncertainty | R&S® NRP6A | @ -23.3 dBm after external zeroing |  |  | 0.081264 | 0.081218 |  | 0.081 | 0.081
Expanded level measurement uncertainty | R&S® NRP6A | @ -30 dBm after external zeroing |  |  | 0.082233 | 0.08202 |  | 0.081007 | 0.081006
Expanded level measurement uncertainty | R&S® NRP6A | @ -35 dBm after external zeroing |  |  | 0.084899 | 0.084226 |  | 0.081067 | 0.081064
Expanded level measurement uncertainty | R&S® NRP6A | @ -42.7 dBm after external zeroing |  |  | 0.10392 | 0.099963 |  | 0.083284 | 0.08319
Expanded level measurement uncertainty | R&S® NRP6A | @ -50 dBm after external zeroing |  |  | 0.20311 | 0.18188 |  | 0.1311 | 0.12937
Expanded level measurement uncertainty | R&S® NRP6A | @ -51 dBm after external zeroing |  |  | 0.23433 | 0.20762 |  | 0.15267 | 0.15031
Expanded level measurement uncertainty | R&S® NRP6A | @ -55 dBm after external zeroing |  |  | 0.45909 | 0.39231 |  | 0.32853 | 0.32168
Expanded level measurement uncertainty | R&S® NRP6A | @ -58 dBm after external zeroing |  |  | 0.8137 | 0.68146 |  | 0.62016 | 0.6059
Expanded level measurement uncertainty | R&S® NRP6A | @ -63 dBm after external zeroing |  |  | 2.156 | 1.7509 |  | 1.7202 | 1.6718
Absolute level measurement uncertainty | R&S® NRP-Z91 | between -67 dBm and -19 dBm (+15 °C to +35 °C) | 0.081 | 0.081 | [6 empty] | Zero offset absolute (pW) | 470 | Noise absolute (pW) | 65
Absolute level measurement uncertainty | R&S® NRP-Z91 | between -19 dBm and +1 dBm (+15 °C to +35 °C) | 0.077 | 0.077
| R&S® NRP-Z91 | @ -4 dBm after external zeroing | 5.1272e-06 | 5.1272e-06 | [6 empty] | Signal level (dBm) | -4
| R&S® NRP-Z91 | @ -4 dBm | 7.0908e-07 | 0
| R&S® NRP-Z91 | @ -12 dBm after external zeroing | 3.235e-05 | 3.235e-05 | [6 empty] | Signal level (dBm) | -12
| R&S® NRP-Z91 | @ -12 dBm | 4.474e-06 | 0
Zero offset | R&S® NRP-Z91 | @ -23.3 dBm after external zeroing | 0.00043638 | 0.00043638 | [6 empty] | Signal level (dBm) | -23.3
Measurement noise | R&S® NRP-Z91 | @ -23.3 dBm | 6.0352e-05 | 0
Zero offset | R&S® NRP-Z91 | @ -30 dBm after external zeroing | 0.0020407 | 0.0020407 | [6 empty] | Signal level (dBm) | -30
Measurement noise | R&S® NRP-Z91 | @ -30 dBm | 0.00028228 | 0
Zero offset | R&S® NRP-Z91 | @ -35 dBm after external zeroing | 0.00645 | 0.00645 | [6 empty] | Signal level (dBm) | -35
Measurement noise | R&S® NRP-Z91 | @ -35 dBm | 0.00089259 | 0
Zero offset | R&S® NRP-Z91 | @ -42.7 dBm after external zeroing | 0.037843 | 0.037843 | [6 empty] | Signal level (dBm) | -42.7
Measurement noise | R&S® NRP-Z91 | @ -42.7 dBm | 0.0052533 | 0
Zero offset | R&S® NRP-Z91 | @ -50 dBm after external zeroing | 0.19947 | 0.19947 | [6 empty] | Signal level (dBm) | -50
Measurement noise | R&S® NRP-Z91 | @ -50 dBm | 0.028138 | 0
Zero offset | R&S® NRP-Z91 | @ -51 dBm after external zeroing | 0.24965 | 0.24965 | [6 empty] | Signal level (dBm) | -51
Measurement noise | R&S® NRP-Z91 | @ -51 dBm | 0.035394 | 0
Zero offset | R&S® NRP-Z91 | @ -55 dBm after external zeroing | 0.60179 | 0.60179 | [6 empty] | Signal level (dBm) | -55
Measurement noise | R&S® NRP-Z91 | @ -55 dBm | 0.088363 | 0
Zero offset | R&S® NRP-Z91 | @ -58 dBm after external zeroing | 1.1279 | 1.1279 | [6 empty] | Signal level (dBm) | -58
Measurement noise | R&S® NRP-Z91 | @ -58 dBm | 0.17456 | 0
Zero offset | R&S® NRP-Z91 | @ -63 dBm after external zeroing | 2.873 | 2.873 | [6 empty] | Signal level (dBm) | -63
Measurement noise | R&S® NRP-Z91 | @ -63 dBm | 0.5296 | 0
Expanded level measurement uncertainty | R&S® NRP-Z91 | @ -4 dBm after external zeroing |  |  | 0.077006 | 0.077005 |  | 0.077 | 0.077
Expanded level measurement uncertainty | R&S® NRP-Z91 | @ -12 dBm after external zeroing |  |  | 0.077037 | 0.077032 |  | 0.077 | 0.077
Expanded level measurement uncertainty | R&S® NRP-Z91 | @ -23.3 dBm after external zeroing |  |  | 0.081497 | 0.081436 |  | 0.081001 | 0.081001
Expanded level measurement uncertainty | R&S® NRP-Z91 | @ -30 dBm after external zeroing |  |  | 0.083323 | 0.083041 |  | 0.081026 | 0.081026
Expanded level measurement uncertainty | R&S® NRP-Z91 | @ -35 dBm after external zeroing |  |  | 0.088343 | 0.08745 |  | 0.081261 | 0.081256
Expanded level measurement uncertainty | R&S® NRP-Z91 | @ -42.7 dBm after external zeroing |  |  | 0.1241 | 0.11884 |  | 0.089558 | 0.089404
Expanded level measurement uncertainty | R&S® NRP-Z91 | @ -50 dBm after external zeroing |  |  | 0.3086 | 0.28047 |  | 0.21712 | 0.21529
Expanded level measurement uncertainty | R&S® NRP-Z91 | @ -51 dBm after external zeroing |  |  | 0.36605 | 0.33065 |  | 0.26484 | 0.26247
Expanded level measurement uncertainty | R&S® NRP-Z91 | @ -55 dBm after external zeroing |  |  | 0.77115 | 0.68279 |  | 0.61361 | 0.60722
Expanded level measurement uncertainty | R&S® NRP-Z91 | @ -58 dBm after external zeroing |  |  | 1.3835 | 1.2089 |  | 1.1442 | 1.1308
Expanded level measurement uncertainty | R&S® NRP-Z91 | @ -63 dBm after external zeroing |  |  | 3.4836 | 2.954 |  | 2.9226 | 2.8742
Absolute level measurement uncertainty | R&S® NRP-Z92 | between -57 dBm and +20 dBm (+15 °C to +35 °C) | 0.106 | 0.106 | [6 empty] | Zero offset absolute (pW) | 1300 | Noise absolute (pW) | 800
Zero offset | R&S® NRP-Z92 | @ -4 dBm after external zeroing | 1.4182e-05 | 1.4182e-05 | [6 empty] | Signal level (dBm) | -4
Measurement noise | R&S® NRP-Z92 | @ -4 dBm | 8.7272e-06 | 0
Zero offset | R&S® NRP-Z92 | @ -12 dBm after external zeroing | 8.9479e-05 | 8.9479e-05 | [6 empty] | Signal level (dBm) | -12
Measurement noise | R&S® NRP-Z92 | @ -12 dBm | 5.5064e-05 | 0
Zero offset | R&S® NRP-Z92 | @ -23.3 dBm after external zeroing | 0.0012069 | 0.0012069 | [6 empty] | Signal level (dBm) | -23.3
Measurement noise | R&S® NRP-Z92 | @ -23.3 dBm | 0.00074274 | 0
Zero offset | R&S® NRP-Z92 | @ -30 dBm after external zeroing | 0.0056422 | 0.0056422 | [6 empty] | Signal level (dBm) | -30
Measurement noise | R&S® NRP-Z92 | @ -30 dBm | 0.003473 | 0
Zero offset | R&S® NRP-Z92 | @ -35 dBm after external zeroing | 0.017817 | 0.017817 | [6 empty] | Signal level (dBm) | -35
Measurement noise | R&S® NRP-Z92 | @ -35 dBm | 0.010973 | 0
Zero offset | R&S® NRP-Z92 | @ -42.7 dBm after external zeroing | 0.10388 | 0.10388 | [6 empty] | Signal level (dBm) | -42.7
Measurement noise | R&S® NRP-Z92 | @ -42.7 dBm | 0.064218 | 0
Zero offset | R&S® NRP-Z92 | @ -50 dBm after external zeroing | 0.53078 | 0.53078 | [6 empty] | Signal level (dBm) | -50
Measurement noise | R&S® NRP-Z92 | @ -50 dBm | 0.33424 | 0
Zero offset | R&S® NRP-Z92 | @ -51 dBm after external zeroing | 0.65826 | 0.65826 | [6 empty] | Signal level (dBm) | -51
Measurement noise | R&S® NRP-Z92 | @ -51 dBm | 0.41675 | 0
Zero offset | R&S® NRP-Z92 | @ -55 dBm after external zeroing | 1.4956 | 1.4956 | [6 empty] | Signal level (dBm) | -55
Measurement noise | R&S® NRP-Z92 | @ -55 dBm | 0.97945 | 0
Zero offset | R&S® NRP-Z92 | @ -58 dBm after external zeroing | 2.6013 | 2.6013 | [6 empty] | Signal level (dBm) | -58
Measurement noise | R&S® NRP-Z92 | @ -58 dBm | 1.7747 | 0
Zero offset | R&S® NRP-Z92 | @ -63 dBm after external zeroing | 5.5556 | 5.5556 | [6 empty] | Signal level (dBm) | -63
Measurement noise | R&S® NRP-Z92 | @ -63 dBm | 4.1434 | 0
Expanded level measurement uncertainty | R&S® NRP-Z92 | @ -4 dBm after external zeroing |  |  | 0.10602 | 0.10601 |  | 0.106 | 0.106
Expanded level measurement uncertainty | R&S® NRP-Z92 | @ -12 dBm after external zeroing |  |  | 0.10614 | 0.10609 |  | 0.106 | 0.106
Expanded level measurement uncertainty | R&S® NRP-Z92 | @ -23.3 dBm after external zeroing |  |  | 0.10795 | 0.10721 |  | 0.10601 | 0.10601
Expanded level measurement uncertainty | R&S® NRP-Z92 | @ -30 dBm after external zeroing |  |  | 0.11512 | 0.11164 |  | 0.10621 | 0.10615
| R&S® NRP-Z92 | @ -35 dBm after external zeroing |  |  | 0.13479 | 0.12382 |  | 0.10805 | 0.10749
Expanded level measurement uncertainty | R&S® NRP-Z92 | @ -42.7 dBm after external zeroing |  |  | 0.2741 | 0.20988 |  | 0.16171 | 0.14841
Expanded level measurement uncertainty | R&S® NRP-Z92 | @ -50 dBm after external zeroing |  |  | 0.97102 | 0.63678 |  | 0.63615 | 0.54127
Expanded level measurement uncertainty | R&S® NRP-Z92 | @ -51 dBm after external zeroing |  |  | 1.181 | 0.76426 |  | 0.78627 | 0.66674
Expanded level measurement uncertainty | R&S® NRP-Z92 | @ -55 dBm after external zeroing |  |  | 2.581 | 1.6016 |  | 1.7909 | 1.4993
Expanded level measurement uncertainty | R&S® NRP-Z92 | @ -58 dBm after external zeroing |  |  | 4.482 | 2.7073 |  | 3.1508 | 2.6035
Expanded level measurement uncertainty | R&S® NRP-Z92 | @ -63 dBm after external zeroing |  |  | 9.805 | 5.6616 |  | 6.9313 | 5.5566
Absolute level uncertainty at 64 MHz | R&S® FSV | +20 °C to +30 °C | 0.2 | 0.2
Frequency response re. 64 MHz | R&S® FSV | 10 MHz ≤ f < 3.6 GHz | 0.3 | 0.3 |  | Total level uncertainty |  | 0.28 | 0.28
Sensitivity result | R&S® NRP6A |  |  |  | 1.932 | 2.0322 | 2.0322 | 0.84216 | 0.85397 | 0.85397
Sensitivity result | R&S® NRP-Z91 |  |  |  | 1.933 | 2.0333 | 2.0333 | 0.84226 | 0.85407 | 0.85407
Sensitivity result | R&S® NRP-Z92 |  |  |  | 1.9616 | 2.0651 | 2.0651 | 0.84556 | 0.85776 | 0.85776
Sensitivity result | R&S® FSV |  |  |  | 2.35 | 2.45 | 2.45 | 0.88369 | 0.88369 | 0.88369
Adjacent channel selectivity result 1 | R&S® NRP6A |  |  |  | 0.28888 | 0.31555 | 0.31555 | 0.17075 | 0.17282 | 0.17282
Adjacent channel selectivity result 1 | R&S® NRP-Z91 |  |  |  | 0.41215 | 0.44748 | 0.44748 | 0.27468 | 0.27695 | 0.27695
Adjacent channel selectivity result 1 | R&S® NRP-Z92 |  |  |  | 0.87221 | 1.2882 | 1.2882 | 0.67512 | 0.79338 | 0.79338
Adjacent channel selectivity result 1 | R&S® FSV |  |  |  | 1 | 1 | 1 | 0.39598 | 0.39598 | 0.39598
Adjacent channel selectivity result 2 | R&S® NRP6A | C level increased by 30 dB for meas. |  |  | 1.7633 | 1.7635 | 1.7635 | 0.83254 | 0.83254 | 0.83254
Adjacent channel selectivity result 2 | R&S® NRP-Z91 | C level increased by 30 dB for meas. |  |  | 1.7645 | 1.7648 | 1.7648 | 0.83254 | 0.83254 | 0.83254
Adjacent channel selectivity result 2 | R&S® NRP-Z92 | C level increased by 30 dB for meas. |  |  | 1.8196 | 1.8223 | 1.8223 | 0.83912 | 0.83816 | 0.83912
Adjacent channel selectivity result 2 | R&S® FSV |  |  |  | 1 | 1 | 1 | 0.39598 | 0.39598 | 0.39598
Adjacent channel selectivity result 3 DVB-T | R&S® NRP6A | C level increased by 30 dB for meas. |  |  | 1.7859 | 1.7822 | 1.7859 | 0.83277 | 0.83276 | 0.83277
Adjacent channel selectivity result 3 DVB-T | R&S® NRP-Z91 | C level increased by 30 dB for meas. |  |  | 1.8071 | 1.8022 | 1.8071 | 0.83342 | 0.8334 | 0.83342
Adjacent channel selectivity result 3 DVB-T | R&S® NRP-Z92 | C level increased by 30 dB for meas. |  |  | 1.9857 | 1.925 | 1.9857 | 0.84701 | 0.84457 | 0.84701
Adjacent channel selectivity result 3 DVB-T | R&S® FSV |  |  |  | 1 | 1 | 1 | 0.39598 | 0.39598 | 0.39598
Adjacent channel selectivity result 3 DVB-T2 | R&S® NRP6A | C level increased by 30 dB for meas. |  |  | 1.7881 | 1.7849 | 1.7881 | 0.83277 | 0.83276 | 0.83277
Adjacent channel selectivity result 3 DVB-T2 | R&S® NRP-Z91 | C level increased by 30 dB for meas. |  |  | 1.8115 | 1.8072 | 1.8115 | 0.83344 | 0.83342 | 0.83344
Adjacent channel selectivity result 3 DVB-T2 | R&S® NRP-Z92 | C level increased by 30 dB for meas. |  |  | 1.9979 | 1.9447 | 1.9979 | 0.84717 | 0.84481 | 0.84717
Adjacent channel selectivity result 3 DVB-T2 | R&S® FSV |  |  |  | 1 | 1 | 1 | 0.39598 | 0.39598 | 0.39598
Adjacent channel selectivity result 4 and 5 | R&S® NRP6A |  |  |  | 0.47454 | 0.54031 | 0.54031 | 0.33172 | 0.33837 | 0.33837
Adjacent channel selectivity result 4 and 5 | R&S® NRP-Z91 |  |  |  | 0.76611 | 0.85259 | 0.85259 | 0.6126 | 0.61894 | 0.61894
Adjacent channel selectivity result 4 and 5 | R&S® NRP-Z92 | C level increased by 30 dB for meas. |  |  | 1.8268 | 1.8268 | 1.8268 | 0.83818 | 0.83818 | 0.83818
Adjacent channel selectivity result 4 and 5 | R&S® FSV |  |  |  | 1 | 1 | 1 | 0.39598 | 0.39598 | 0.39598
Blocking result | R&S® NRP6A |  |  |  | 1.7643 | 1.7643 | 1.7643 | 0.69312 | 0.69312 | 0.69312
Blocking result | R&S® NRP-Z91 |  |  |  | 1.6 | 1.6 | 1.6 | 0.68 | 0.68 | 0.68
Blocking result | R&S® NRP-Z92 |  |  |  | 1.8268 | 1.8268 | 1.8268 | 0.70255 | 0.70255 | 0.70255
Blocking result | R&S® FSV |  |  |  | 1 | 1 | 1 | 0.39598 | 0.39598 | 0.39598
Overloading result 1 | R&S® NRP6A |  |  |  | 0.41792 | 0.51725 | 0.51725 | 0.18963 | 0.23656 | 0.23656
Overloading result 1 | R&S® NRP-Z91 |  |  |  | 0.41537 | 0.51449 | 0.51449 | 0.18717 | 0.23459 | 0.23459
Overloading result 1 | R&S® NRP-Z92 |  |  |  | 0.49088 | 0.57996 | 0.57996 | 0.2131 | 0.25552 | 0.25552
Overloading result 1 | R&S® FSV |  |  |  | 1.25 | 1.35 | 1.35 | 0.42344 | 0.44643 | 0.44643
Overloading result 2 and 3 | R&S® NRP6A |  |  |  | 0.4179 | 0.51723 | 0.51723 | 0.18963 | 0.23656 | 0.23656
Overloading result 2 and 3 | R&S® NRP-Z91 |  |  |  | 0.41535 | 0.51446 | 0.51446 | 0.18717 | 0.23459 | 0.23459
Overloading result 2 and 3 | R&S® NRP-Z92 |  |  |  | 0.4908 | 0.57984 | 0.57984 | 0.2131 | 0.25552 | 0.25552
Overloading result 2 and 3 | R&S® FSV |  |  |  | 1.25 | 1.35 | 1.35 | 0.42344 | 0.44643 | 0.44643